## Wireless Node Config
LORD Sensing Wireless Node EEPROM Map
EEPROM Address |  | Option | Type | Valid Values | Description | Min FW Version | Category
decimal | hex
0 | 0x0 | Current Datalogging Page | uint16 |  | The flash memory page that will be used for the next datalogging session. If current page offset >= 264, multiply by 2 and add 1, else multiply by 2. Do Not Change. | 1 | logging
2 | 0x2 | Current Datalogging Page Offset | uint16 |  | The byte offset for the next datalogging session. If current page offset >= 264, subtract 264. Do Not Change. | 1 | logging
4 | 0x4 | # of Datalogging Sessions Stored | uint16 |  | The number of datalogging sessions (triggers) stored in flash memory. Do Not Change. | 1 | logging
6 | 0x6 | (LEGACY) Sensor Trigger Value | uint16 |  |  | 1 | -
8 | 0x8 | (LEGACY) Sensor Trigger Channel | uint16 |  |  | 1 | -
10 | 0xA | (LEGACY) Sensor Trigger Type | uint16 |  |  | 1 | -
12 | 0xC | Active Channel Mask | uint16 | see notes | The channels that are enabled/disabled for sampling. | 1 | -
14 | 0xE | Datalogging Sample Rate | uint16 | see notes | The sample rate for Armed Datalogging sampling. | 1 | logging
16 | 0x10 | # of Sweeps | uint16 |  | The number of sweeps (samples per data set) to sample for. | 1 | -
18 | 0x12 | Default Mode | uint16 | see notes | The default mode to perform on startup and user inactivity timeout. | 1 | -
20 | 0x14 | Serial ID | uint32 |  | The serial id for the device. Combine with the model number and model option for the full Serial Number. | 1 | -
22 | 0x16
24 | 0x18 | Sampling Mode | uint16 | see notes | The sampling mode that the device was last placed in (needs to be user updated). | 1 | -
26 | 0x1A | Hardware Offset 1 | uint16 |  | The hardware offset for balancing the sensors output and compensating for sensor offset. (1) | 1 | -
28 | 0x1C | Hardware Offset 2 | uint16 |  | The hardware offset for balancing the sensors output and compensating for sensor offset. (2) | 1 | -
30 | 0x1E | Hardware Offset 3 | uint16 |  | The hardware offset for balancing the sensors output and compensating for sensor offset. (3) | 1 | -
32 | 0x20 | Hardware Offset 4 | uint16 |  | The hardware offset for balancing the sensors output and compensating for sensor offset. (4) | 1 | -
34 | 0x22 | Sampling Delay | uint16 | see notes | The delay to use before sampling, for low-impedance sensors with slow rise/settling times. | 1 | -
36 | 0x24 | TDMA Address | uint16 |  | The TDMA Address for Synchronized Sampling. | 7 | -
38 | 0x26 | Data Collection Method | uint16 | see notes | The Data collection method for Non-Sync and Sync Sampling. | 1 | logging
40 | 0x28 | # of Buffered Packets | uint16 | [1-48] | The payload packet size for Non-Sync sampling. | 7 | -
44 | 0x2C | # of Retransmission Attempts | uint16 |  | The number of retransmission attempts during Lossless sampling before entering a low-power state. | 7 | -
46 | 0x2E | Model Number | uint16 |  | The model number of the device. | 8 | -
48 | 0x30 | Model Option | uint16 |  | The model option of the device. | 8 | -
50 | 0x32 | Node Address | uint16 | [1-65534] | The address of the device. | 1 | -
56 | 0x38 | (LEGACY) Calibration | uint16 |  | (LEGACY) Used in performing an onboard calibration of strain sensors. | 1 | -
66 | 0x42 | Set to Idle Interval | uint16 | see notes | How often the Node wakes up and listens for a 'set to idle' command. | 1 | -
70 | 0x46 | User Inactivity Timeout | uint16 | [5-65535] | The time, in seconds, before a Node enters sleep mode, if there is no user activity. | 1 | -
72 | 0x48 | Sample Rate | uint16 | see notes | The sample rate for Non-Sync and Sync Sampling. | 1 | -
74 | 0x4A | Auto Balance Interval | uint16 |  |  | 1 | -
76 | 0x4C | Data Format | uint16 | see notes | The format that data is transferred in for Non-Sync and Sync Sampling modes. | 1 | -
78 | 0x4E | Sniff Duration | uint16 | [1-50] | The amount of time the Node listens for a Set to Idle packet for each check (see Set to Idle Interval). | 1 | -
80 | 0x50 | (LEGACY) Manual Trigger Config | uint16 |  |  | 1 | -
82 | 0x52 | Clock Timer Control | uint16 | [0-15] | Internal Use Only. Do Not Change. | 1 | -
88 | 0x58 | Radio Broadcast Filter | uint16 |  | Prevents the node from hearing "broadcast" packets in certain modes. 0 = disabled, 1 = sleep mode | 7 | -
90 | 0x5A | Radio Frequency | uint16 | see notes | The radio frequency of the device. Needs to be on the same frequency as the BaseStation to communicate with it. | 1 | -
94 | 0x5E | Transmit Power Level | uint16 | see notes | The transmit output power of the radio. | 1 | -
96 | 0x60 | Return Node Address | uint16 | 4660 | Internal Use. Do Not Change. | 1 | -
100 | 0x64 | Unlimited Sampling | uint16 | [0, 1] | Enables (1) or Disables (0) unlimited duration of Non-Sync and Sync Sampling. | 1 | -
102 | 0x66 | Unlimited Armed Datalogging | uint16 | [0, 1] | Enables (1) or Disables (0) unlimited duration of Armed Datalogging Sampling. | 1 | logging
108 | 0x6C | Firmware Version (1) | uint16 | see notes | The firmware version of the device (Part 1). | 4 | -
110 | 0x6E | Firmware Version (2) | uint16 | see notes | The firmware version of the device (Part 2). | 4 | -
112 | 0x70 | (LEGACY) Model Number | uint16 |  | (LEGACY, Replaced by EEPROMs 46 + 48) The model number of the device. | 1 | -
114 | 0x72 | (LEGACY) Serial ID | uint16 |  | (LEGACY, Replaced by EEPROM 20) The serial id for the device. | 1 | -
116 | 0x74 | Max Memory | uint16 |  | The maximum number of pages available on the flash memory. | 1 | logging
118 | 0x76 | Radio ID | uint16 |  | The id of the radio on the device. | 1 | -
120 | 0x78 | Microcontroller ID | uint16 | see notes | The id of the microcontroller on the device. | 1 | -
122 | 0x7A | Firmware Arch Version | uint16 |  | The architecture version of the device used to determine which version of firmware can be applied to it. | 3 | -
124 | 0x7C | ASPP Version LXRS | uint16 |  | The ASPP version (msb.lsb) the device supports when using the LXRS communication protocol. | 10.33793
126 | 0x7E | ASPP Version LXRS+ | uint16 |  | The ASPP version (msb.lsb) the device supports when using the LXRS+ communication protocol. | 11.0 | -
128 | 0x80 | Hardware Gain 1 | uint16 |  | The hardware gain applied to the Node. (1) | 1 | -
130 | 0x82 | Hardware Gain 2 | uint16 |  | The hardware gain applied to the Node. (2) | 1 | -
130 | 0x82 | Filter (1) | uint16 |  | The first filter value for nodes with filter functionality (ie. TC-Link) | 1 | -
132 | 0x84 | Hardware Gain 3 | uint16 |  | The hardware gain applied to the Node. (3) | 1 | -
134 | 0x86 | Hardware Gain 4 | uint16 |  | The hardware gain applied to the Node. (4) | 1 | -
134 | 0x86 | Filter (2) | uint16 |  | The second filter value for nodes with 2 filter functionality (ie. ENV-Link Pro) | 1 | -
136 | 0x88 | Flash ID | uint16 | see notes | Identifier for the flash chip on the Node. | 10
138 | 0x8A | Log Mode | uint16 |  | How the flash logging works. 0 = FIFO (circular buffer). 1 = Log to End | 10 | logging
140 | 0x8C | Log Valid | uint16 |  | (Internal Use) 0 if log start and end are invalid, 1 if they are valid | 10 | logging
142 | 0x8E | Log Start Address | uint32 |  | (Internal Use) Flash address of first logged data block. | 10 | logging
144 | 0x90
146 | 0x92 | Log End Address | uint32 |  | (Internal Use) Flash address of last logged data block. | 10 | logging
148 | 0x94
150 | 0x96 | Action ID - Ch1 | uint16 | see notes | The action id (unit and equation) for channel 1. | 1 | cal coefficients
152 | 0x98 | Slope - Ch1 | float |  | The slope for channel 1. | 1 | cal coefficients
154 | 0x9A
156 | 0x9C | Offset - Ch1 | float |  | The offset for channel 1. | 1 | cal coefficients
158 | 0x9E
160 | 0xA0 | Action ID - Ch2 | uint16 | see notes | The action id (unit and equation) for channel 2. | 1 | cal coefficients
162 | 0xA2 | Slope - Ch2 | float |  | The slope for channel 2. | 1 | cal coefficients
164 | 0xA4
166 | 0xA6 | Offset - Ch2 | float |  | The offset for channel 2. | 1 | cal coefficients
168 | 0xA8
170 | 0xAA | Action ID - Ch3 | uint16 | see notes | The action id (unit and equation) for channel 3. | 1 | cal coefficients
172 | 0xAC | Slope - Ch3 | float |  | The slope for channel 3. | 1 | cal coefficients
174 | 0xAE
176 | 0xB0 | Offset - Ch3 | float |  | The offset for channel 3. | 1 | cal coefficients
178 | 0xB2
180 | 0xB4 | Action ID - Ch4 | uint16 | see notes | The action id (unit and equation) for channel 4. | 1 | cal coefficients
182 | 0xB6 | Slope - Ch4 | float |  | The slope for channel 4. | 1 | cal coefficients
184 | 0xB8
186 | 0xBA | Offset - Ch4 | float |  | The offset for channel 4. | 1 | cal coefficients
188 | 0xBC
190 | 0xBE | Action ID - Ch5 | uint16 | see notes | The action id (unit and equation) for channel 5. | 1 | cal coefficients
192 | 0xC0 | Slope - Ch5 | float |  | The slope for channel 5. | 1 | cal coefficients
194 | 0xC2
196 | 0xC4 | Offset - Ch5 | float |  | The offset for channel 5. | 1 | cal coefficients
198 | 0xC6
200 | 0xC8 | Action ID - Ch6 | uint16 | see notes | The action id (unit and equation) for channel 6. | 1 | cal coefficients
202 | 0xCA | Slope - Ch6 | float |  | The slope for channel 6. | 1 | cal coefficients
204 | 0xCC
206 | 0xCE | Offset - Ch6 | float |  | The offset for channel 6. | 1 | cal coefficients
208 | 0xD0
210 | 0xD2 | Action ID - Ch7 | uint16 | see notes | The action id (unit and equation) for channel 7. | 1 | cal coefficients
212 | 0xD4 | Slope - Ch7 | float |  | The slope for channel 7. | 1 | cal coefficients
214 | 0xD6
216 | 0xD8 | Offset - Ch7 | float |  | The offset for channel 7. | 1 | cal coefficients
218 | 0xDA
220 | 0xDC | Action ID - Ch8 | uint16 | see notes | The action id (unit and equation) for channel 8. | 1 | cal coefficients
222 | 0xDE | Slope - Ch8 | float |  | The slope for channel 8. | 1 | cal coefficients
224 | 0xE0
226 | 0xE2 | Offset - Ch8 | float |  | The offset for channel 8. | 1 | cal coefficients
228 | 0xE4
230 | 0xE6 | (LEGACY) Bootload Page Address | uint16 |  | Internal Use Only. Do Not Change. | 1 | -
232 | 0xE8 | (LEGACY) Bootload Enable Flag | uint16 |  | Internal Use Only. Do Not Change. | 1 | -
234 | 0xEA | (LEGACY) Bootload Checksum | uint16 |  | Internal Use Only. Do Not Change. | 1 | -
236 | 0xEC | (LEGACY) Bootload Version | uint16 |  | Internal Use Only. Do Not Change. | 1 | -
238 | 0xEE | (LEGACY) Bootload Options | uint16 |  | Internal Use Only. Do Not Change. | 1 | -
240 | 0xF0 | (LEGACY) Bootload Counter | uint16 |  | Internal Use Only. Do Not Change. | 1 | -
246 | 0xF6 | (LEGACY) Reset Counter | uint16 |  | Records the number of times the device has been reset (both soft reset and power cycling). | 1 | -
248 | 0xF8 | Brown Out Reset Enable | uint16 |  | Enables (1) or Disables (0) the brown out reset function of the microcontroller during sleep. If a value of 0 is used, the brown out reset is disabled during sleep. This will result in the lowest sleep power consumption, but it may be difficult to turn off or reset a sleeping node. (The power switch may need to be placed in the off position for up to 30 seconds). If any non zero value is written, the brown out reset will be enabled during sleep. This will add approximately 40 uA to the overall sleep current, but the node will turn off immediately when the switch is powered off. | 1 | -
250 | 0xFA | Cycle Power | uint16 | [1, 2] | Used to cycle the power on the device. (1-power, 2-radio) | 1 | -
254 | 0xFE | Startup Mode Override (SUMO) Settings | uint16 |  | A value of 100 will disable SUMO (this may cause the device to be unusable, so do this with caution). | 4.98 | -
258 | 0x102 | Radio Options | uint16 |  | The radio options of the Node. | 1 | -
260 | 0x104 | Lost Beacon Timeout | uint16 |  | Set a value in minutes between 2 and 600. If a node is in synchronized sampling mode and loses the beacon for a length of time greater than this value, then the node drops into a sleep mode. The node will re-enter synchronized sampling within 2 minutes of the beacon re-appearing. This feature is disabled when the value entered in the field is greater than 600. | 1 | -
262 | 0x106 | Sync Sampling Settings | uint16 | see notes | The Sync Sampling settings (continuous, burst, etc.). | 7 | -
264 | 0x108 | Transmissions Per Group | uint16 |  | The number of transmissions per group (for Sync Sampling). | 7 | -
266 | 0x10A | Group Size | uint16 |  | The size of each transmission group (for Sync Sampling). | 7 | -
268 | 0x10C | Time Between Sessions | uint16 |  | The time between each burst for Sync Sampling. | 7 | -
270 | 0x10E | Armed Datalogging & Beacon Timeout | uint16 |  | Sets the number of seconds that the node will wait in ARMED mode before being triggered for datalogging. This timeout also determines the number of seconds the node will wait in ARMED mode before receiving a valid beacon packet when entering a synchronized sampling session. When the value within this field is out of range it will use a default timeout of 10 seconds. | 7 | -
272 | 0x110 | Node Retransmission | uint16 | see notes | Enables and disables retransmission (lossless) for Sync Sampling. | 7 | -
274 | 0x112 | Max Retransmissions per Packet | uint16 |  | The maximum allowed retransmissions per packet (for Sync Sampling). | 7 | -
276 | 0x114 | Max Retransmissions per Burst | uint16 |  | The maximum allowed retransmissions per burst (for Sync Sampling). | 7 | -
278 | 0x116 | Lossless Buffer Size | uint16 |  | The value stored, multiplied by 10, equals the total number of Bytes that may be held in the lossless buffer. | 7 | -
280 | 0x118 | Region Code | uint16 | see notes | The region the device will be operated in. | 8 | -
282 | 0x11A | Data Mode | uint16 |  | The data mode to use, as a bitmask. (raw = bit 1, math = bit 2)
284 | 0x11C | Derived Data Rate | uint16 |  | The sample rate for all of the Derived Channel Data.
286 | 0x11E | Derived RMS Mask | uint16 |  | The channel mask for the Derived RMS Data.
288 | 0x120 | Derived Peak to Peak Mask | uint16 |  | The channel mask for the Derived Peak to Peak Data.
290 | 0x122 | Derived IPS Mask | uint16 |  | The channel mask for the Derived IPS Data.
292 | 0x124 | Data Packet Format (internal use only) | uint16 |  | The format of the data in data packets that will be sent out when sampling (used internal by the Node).
294 | 0x126 | Derived Crest Factor Mask | uint16 |  | The channel mask for the Derived Crest Factor Data.
296 | 0x128 | Derived Mean Mask | uint16 |  | The channel mask for the Derived Mean (Average) Data.
300 | 0x12C | Diagnostic Packet control | uint16 |  | Enables or disables the Diagnostic Packet. A value of 0 disables the packet. Any other value is the interval in seconds. | 10.3096
306 | 0x132 | Thermocouple Type | uint16 |  | The thermocouple type for nodes with thermocouple functionality (ie. TC-Link) | 1 | -
328 | 0x148 | Event Sample Rate | uint16 |  | The sample rate to collect data for after the event is triggered. | 1 | event
330 | 0x14A | Event Mask | uint16 |  | Enables/Disables specific Events. | 1 | event
332 | 0x14C | Event Pre Duration | uint16 |  | The amount of time before the event occurred to send data (in seconds). | 1 | event
334 | 0x14E | Event Post Duration | uint16 |  | The amount of time after the event occurred to send data (in seconds). | 1 | event
336 | 0x150 | Event Source - Event 1 | uint16 |  | The source for event 1. | 1 | event
338 | 0x152 | Event Operation -Event 1 | uint16 |  | The operation for event 1 . | 1 | event
340 | 0x154 | Event Value 1 - Event 1 | uint16 |  | Value 1 of event 1. | 1 | event
342 | 0x156 | Event Value 2 - Event 1 | uint16 |  | Value 2 of event 1. | 1 | event
344 | 0x158 | Event Source - Event 2 | uint16 |  | The source for event 2. | 1 | event
346 | 0x15A | Event Operation -Event 2 | uint16 |  | The operation for event 2. | 1 | event
348 | 0x15C | Event Value 1 - Event 2 | uint16 |  | Value 1 of event 2. | 1 | event
350 | 0x15E | Event Value 2 - Event 2 | uint16 |  | Value 2 of event 2. | 1 | event
352 | 0x160 | Event Source - Event 3 | uint16 |  | The source for event 3. | 1 | event
354 | 0x162 | Event Operation -Event 3 | uint16 |  | The operation for event 3. | 1 | event
356 | 0x164 | Event Value 1 - Event 3 | uint16 |  | Value 1 of event 3. | 1 | event
358 | 0x166 | Event Value 2 - Event 3 | uint16 |  | Value 2 of event 3. | 1 | event
360 | 0x168 | Event Source - Event 4 | uint16 |  | The source for event 4. | 1 | event
362 | 0x16A | Event Operation -Event 4 | uint16 |  | The operation for event 4. | 1 | event
364 | 0x16C | Event Value 1 - Event 4 | uint16 |  | Value 1 of event 4. | 1 | event
366 | 0x16E | Event Value 2 - Event 4 | uint16 |  | Value 2 of event 4. | 1 | event
368 | 0x170 | Event Source - Event 5 | uint16 |  | The source for event 5. | 1 | event
370 | 0x172 | Event Operation -Event 5 | uint16 |  | The operation for event 5. | 1 | event
372 | 0x174 | Event Value 1 - Event 5 | uint16 |  | Value 1 of event 5. | 1 | event
374 | 0x176 | Event Value 2 - Event 5 | uint16 |  | Value 2 of event 5. | 1 | event
376 | 0x178 | Event Source - Event 6 | uint16 |  | The source for event 6. | 1 | event
378 | 0x17A | Event Operation -Event 6 | uint16 |  | The operation for event 6. | 1 | event
380 | 0x17C | Event Value 1 - Event 6 | uint16 |  | Value 1 of event 6. | 1 | event
382 | 0x17E | Event Value 2 - Event 6 | uint16 |  | Value 2 of event 6. | 1 | event
384 | 0x180 | Event Source - Event 7 | uint16 |  | The source for event 7. | 1 | event
386 | 0x182 | Event Operation -Event 7 | uint16 |  | The operation for event 7. | 1 | event
388 | 0x184 | Event Value 1 - Event 7 | uint16 |  | Value 1 of event 7. | 1 | event
390 | 0x186 | Event Value 2 - Event 7 | uint16 |  | Value 2 of event 7. | 1 | event
392 | 0x188 | Event Source - Event 8 | uint16 |  | The source for event 8. | 1 | event
394 | 0x18A | Event Operation -Event 8 | uint16 |  | The operation for event 8. | 1 | event
396 | 0x18C | Event Value 1 - Event 8 | uint16 |  | Value 1 of event 8. | 1 | event
398 | 0x18E | Event Value 2 - Event 8 | uint16 |  | Value 2 of event 8. | 1 | event
440 | 0x1B8 | Anti-Aliasing Filter 1 | uint16 |  | The Anti-Aliasing filter for channels on the Node (1). | 10 | -
442 | 0x1BA | Anti-Aliasing Filter 2 | uint16 |  | The Anti-Aliasing filter for channels on the Node (2). | 10 | -
444 | 0x1BC | Anti-Aliasing Filter 3 | uint16 |  | The Anti-Aliasing filter for channels on the Node (3). | 10 | -
446 | 0x1BE | Anti-Aliasing Filter 4 | uint16 |  | The Anti-Aliasing filter for channels on the Node (4). | 10 | -
448 | 0x1C0 | Low-Pass Filter 1 | uint16 |  | The low-pass filter for channels on the Node (1). |  | -
450 | 0x1C2 | Low-Pass Filter 2 | uint16 |  | The low-pass filter for channels on the Node (2). |  | -
452 | 0x1C4 | Low-Pass Filter 3 | uint16 |  | The low-pass filter for channels on the Node (3). |  | -
454 | 0x1C6 | Low-Pass Filter 4 | uint16 |  | The low-pass filter for channels on the Node (4). |  | -
456 | 0x1C8 | Hardware Gain 5 | uint16 |  | The hardware gain applied to the Node. (5) | 10 | -
458 | 0x1CA | Hardware Gain 6 | uint16 |  | The hardware gain applied to the Node. (6) | 10 | -
460 | 0x1CC | Hardware Gain 7 | uint16 |  | The hardware gain applied to the Node. (7) | 10 | -
462 | 0x1CE | Hardware Gain 8 | uint16 |  | The hardware gain applied to the Node. (8) | 10 | -
464 | 0x1D0 | High-Pass Filter 1 | uint16 |  | The high-pass filter for channels on the Node (1). |  | -
466 | 0x1D2 | High-Pass Filter 2 | uint16 |  | The high-pass filter for channels on the Node (2). |  | -
468 | 0x1D4 | High-Pass Filter 3 | uint16 |  | The high-pass filter for channels on the Node (3). |  | -
470 | 0x1D6 | High-Pass Filter 4 | uint16 |  | The high-pass filter for channels on the Node (4). |  | -
480 | 0x1E0 | Min software version required (Major) | uint16 |  | The minimum SensorConnect version required to support the Node. (major) | 10 | -
482 | 0x1E2 | Min software version required (Minor) | uint16 |  | The minimum SensorConnect version required to support the Node. (minor) | 10 | -
504 | 0x1F8 | Hardware Offset 5 | uint16 |  | The hardware offset for balancing the sensors output and compensating for sensor offset. (5) | 1 | -
506 | 0x1FA | Hardware Offset 6 | uint16 |  | The hardware offset for balancing the sensors output and compensating for sensor offset. (6) | 1 | -
508 | 0x1FC | Hardware Offset 7 | uint16 |  | The hardware offset for balancing the sensors output and compensating for sensor offset. (7) | 1 | -
510 | 0x1FE | Hardware Offset 8 | uint16 |  | The hardware offset for balancing the sensors output and compensating for sensor offset. (8) | 1 | -
512 | 0x200 | Custom Polynomial 1 - term 1 | float |  | Term 1 of Polynomial 1 : (x^1) | 1 | polynomial
514 | 0x202
516 | 0x204 | Custom Polynomial 1 - term 2 | float |  | Term 2 of Polynomial 1 : (x^2) | 1 | polynomial
518 | 0x206
520 | 0x208 | Custom Polynomial 1 - term 3 | float |  | Term 3 of Polynomial 1 : (x^3) | 1 | polynomial
522 | 0x20A
524 | 0x20C | Custom Polynomial 1 - term 4 | float |  | Term 4 of Polynomial 1 : (x^4) | 1 | polynomial
526 | 0x20E
528 | 0x210 | Custom Polynomial 1 - term 5 | float |  | Term 5 of Polynomial 1 : (x^5) | 1 | polynomial
530 | 0x212
532 | 0x214 | Custom Polynomial 1 - term 6 | float |  | Term 6 of Polynomial 1 : (x^6) | 1 | polynomial
534 | 0x216
536 | 0x218 | Custom Polynomial 1 - term 7 | float |  | Term 7 of Polynomial 1 : (x^7) | 1 | polynomial
538 | 0x21A
540 | 0x21C | Custom Polynomial 1 - term 8 | float |  | Term 8 of Polynomial 1 : (x^8) | 1 | polynomial
542 | 0x21E
544 | 0x220 | Custom Polynomial 1 - term 9 | float |  | Term 9 of Polynomial 1 : (x^9) | 1 | polynomial
546 | 0x222
548 | 0x224 | Custom Polynomial 1 - term 10 | float |  | Term 10 of Polynomial 1 : (x^10) | 1 | polynomial
550 | 0x226
552 | 0x228 | Custom Polynomial 2 - term 1 | float |  | Term 1 of Polynomial 2 : (x^1) | 1 | polynomial
554 | 0x22A
556 | 0x22C | Custom Polynomial 2 - term 2 | float |  | Term 2 of Polynomial 2 : (x^2) | 1 | polynomial
558 | 0x22E
560 | 0x230 | Custom Polynomial 2 - term 3 | float |  | Term 3 of Polynomial 2 : (x^3) | 1 | polynomial
562 | 0x232
564 | 0x234 | Custom Polynomial 2 - term 4 | float |  | Term 4 of Polynomial 2 : (x^4) | 1 | polynomial
566 | 0x236
568 | 0x238 | Custom Polynomial 2 - term 5 | float |  | Term 5 of Polynomial 2 : (x^5) | 1 | polynomial
570 | 0x23A
572 | 0x23C | Custom Polynomial 2 - term 6 | float |  | Term 6 of Polynomial 2 : (x^6) | 1 | polynomial
574 | 0x23E
576 | 0x240 | Custom Polynomial 2 - term 7 | float |  | Term 7 of Polynomial 2 : (x^7) | 1 | polynomial
578 | 0x242
580 | 0x244 | Custom Polynomial 2 - term 8 | float |  | Term 8 of Polynomial 2 : (x^8) | 1 | polynomial
582 | 0x246
584 | 0x248 | Custom Polynomial 2 - term 9 | float |  | Term 9 of Polynomial 2 : (x^9) | 1 | polynomial
586 | 0x24A
588 | 0x24C | Custom Polynomial 2 - term 10 | float |  | Term 10 of Polynomial 2 : (x^10) | 1 | polynomial
590 | 0x24E
592 | 0x250 | Custom Polynomial 3 - term 1 | float |  | Term 1 of Polynomial 3 : (x^1) | 1 | polynomial
594 | 0x252
596 | 0x254 | Custom Polynomial 3 - term 2 | float |  | Term 2 of Polynomial 3 : (x^2) | 1 | polynomial
598 | 0x256
600 | 0x258 | Custom Polynomial 3 - term 3 | float |  | Term 3 of Polynomial 3 : (x^3) | 1 | polynomial
602 | 0x25A
604 | 0x25C | Custom Polynomial 3 - term 4 | float |  | Term 4 of Polynomial 3 : (x^4) | 1 | polynomial
606 | 0x25E
608 | 0x260 | Custom Polynomial 3 - term 5 | float |  | Term 5 of Polynomial 3 : (x^5) | 1 | polynomial
610 | 0x262
612 | 0x264 | Custom Polynomial 3 - term 6 | float |  | Term 6 of Polynomial 3 : (x^6) | 1 | polynomial
614 | 0x266
616 | 0x268 | Custom Polynomial 3 - term 7 | float |  | Term 7 of Polynomial 3 : (x^7) | 1 | polynomial
618 | 0x26A
620 | 0x26C | Custom Polynomial 3 - term 8 | float |  | Term 8 of Polynomial 3 : (x^8) | 1 | polynomial
622 | 0x26E
624 | 0x270 | Custom Polynomial 3 - term 9 | float |  | Term 9 of Polynomial 3 : (x^9) | 1 | polynomial
626 | 0x272
628 | 0x274 | Custom Polynomial 3 - term 10 | float |  | Term 10 of Polynomial 3 : (x^10) | 1 | polynomial
630 | 0x276
632 | 0x278 | Custom Polynomial - Break Point 1 | float |  | The transition point of when to use polynomial 1 vs polynomial 2 | 1 | polynomial
634 | 0x27A
636 | 0x27C | Custom Polynomial - Break Point 2 | float |  | The transition point of when to use polynomial 2 vs polynomial 3 | 1 | polynomial
638 | 0x27E
640 | 0x280 | Custom Polynomial 1 - Order | uint16 | [1-10] | The order of the polynomial for Polynomial 1. | 1 | polynomial
642 | 0x282 | Custom Polynomial 2 - Order | uint16 | [1-10] | The order of the polynomial for Polynomial 2. | 1 | polynomial
644 | 0x284 | Custom Polynomial 3 - Order | uint16 | [1-10] | The order of the polynomial for Polynomial 3. | 1 | polynomial
682 | 0x2AA | Histogram Bin Range | uint16 |  | The range of each histogram bin. | 7 | shm
684 | 0x2AC | Reset Histogram | uint16 | 1 | Write a 1 to clear the histogram data. | 7 | shm
686 | 0x2AE | Histogram Lower Bound | uint16 |  | The start of the first bin in the histogram. | 7 | shm
688 | 0x2B0 | Peak/Valley Threshold | uint16 |  | The peak/valley threshold. | 7 | shm
700 | 0x2BC | SN-Curve 1 - logA | float |  | logA of the first SN-Curve | 7 | shm
702 | 0x2BE
704 | 0x2C0 | SN-Curve 1 - m | float |  | m of the first SN-Curve | 7 | shm
706 | 0x2C2
708 | 0x2C4 | SN-Curve 2 - logA | float |  | logA of the second SN-Curve | 7 | shm
710 | 0x2C6
712 | 0x2C8 | SN-Curve 2 - m | float |  | m of the second SN-Curve | 7 | shm
714 | 0x2CA
716 | 0x2CC | SN-Curve 3 - logA | float |  | logA of the third SN-Curve | 10 | shm
718 | 0x2CE
720 | 0x2D0 | SN-Curve 3 - m | float |  | m of the third SN-Curve | 10 | shm
722 | 0x2D2
724 | 0x2D4 | Youngs Modulus | float |  | Youngs Modulus | 10 | shm
726 | 0x2D6
728 | 0x2D8 | Poissons Ratio | float |  | Poissons Ratio | 10 | shm
730 | 0x2DA
732 | 0x2DC | Histogram Sample Rate | uint16 |  | Histogram Sample Rate | 10 | shm
734 | 0x2DE | Fatigue - Angle 1 | float |  | The first angle to use for fatigue analysis. | 7 | shm
736 | 0x2E0
738 | 0x2E2 | Fatigue - Angle 2 | float |  | The second angle to use for fatigue analysis. | 7 | shm
740 | 0x2E4
742 | 0x2E6 | Fatigue - Angle 3 | float |  | The third angle to use for fatigue analysis. | 7 | shm
744 | 0x2E8
746 | 0x2EA | Fatigue Mode | uint16 | [0-2] | 0 for standard angle mode, 1 for distributed angle mode, 2 for raw gauge strain mode. | 10 | shm
748 | 0x2EC | Distributed Angle Mode - # Angles | uint16 | [4-16] | The number of angles to use in distributed angle mode. | 10 | shm
750 | 0x2EE | Distributed Angle Mode - Lower Bound | float |  | The lower bound of the angle distribution for distributed angle mode. | 10 | shm
752 | 0x2F0
754 | 0x2F2 | Distributed Angle Mode - Upper Bound | float |  | The upper bound of the angle distribution for distributed angle mode. | 10 | shm
756 | 0x2F4
758 | 0x2F6 | SN-Curve 4 - logA | float |  | logA of the fourth SN-Curve | 7 | shm
760 | 0x2F8
762 | 0x2FA | SN-Curve 4 - m | float |  | m of the fourth SN-Curve | 7 | shm
764 | 0x2FC
766 | 0x2FE | SN-Curve 5 - logA | float |  | logA of the fifth SN-Curve | 7 | shm
768 | 0x300
770 | 0x302 | SN-Curve 5 - m | float |  | m of the fifth SN-Curve | 7 | shm
772 | 0x304
782 | 0x30E | Gauge Factor 1 | float |  | Gauge Factor 1 | 10 | shm
784 | 0x310
786 | 0x312 | Gauge Factor 2 | float |  | Gauge Factor 2 | 10 | shm
788 | 0x314
790 | 0x316 | Gauge Factor 3 | float |  | Gauge Factor 3 | 10 | shm
792 | 0x318
826 | 0x33A | Debug Strain Mode - Enable | uint16 | [0-1] | Whether sending debug strain angle data is enabled (1) or disabled (0). | 10 | shm
834 | 0x342 | Activity Sense - Enable | uint16 | [0-1] | Whether Activity Sense is enabled (1) or disabled (0) | 10 | activity sense
836 | 0x344 | Activity Sense - activity threshold | float |  | Acceleration threshold for activity in G's | 10 | activity sense
838 | 0x346
840 | 0x348 | Activity Sense - inactivity timeout | float |  | Inactivity timeout in seconds | 10 | activity sense
842 | 0x34A
844 | 0x34C | Activity Sense - activity time | float |  | Activity Time in seconds | 10 | activity sense
846 | 0x34E
848 | 0x350 | Activity Sense - inactivity threshold | float |  | Acceleration threshold for inactivity in G's | 10 | activity sense
850 | 0x352
852 | 0x354 | Rainflow Calculation - Enable | uint16 | [0-1] | Whether Rainflow Calculation is enabled (1) or disabled (0). | 10 | shm
856 | 0x358 | Auto Calibration - Maximum Strain | float |  | Units = microstrains. Sets the maximum amount of strain change that can occur while performing an autocalibration. (depracated for fw versions greater than 11.39035) | 11.0 | shm
858 | 0x35A
860 | 0x35C | AutoCalibration - Maximum Bits of Change | float |  | Units = Bits. Sets the maximum change in bits tolerated while performing an autocalibration. | 11.0 | shm
862 | 0x35E
864 | 0x360 | AutoCalibration - number of active gauges | uint16 |  | Number of active strain gauges | 11.0 | shm
1024 | 0x400 | Communication Protocol | uint16 | [0, 1] | The communication protocol the device is currently set for. (0 = LXRS, 1 = LXRS+) | 11.0 | -
1030 | 0x406 | Bootloader version | uint16 | [0-65535] | Bootloader version # (continuous). If changed, the bootloader will rewrite it next time the node boots. | 10 | Bootloader
1032 | 0x408 | Firmware upgrade flag(s) | uint16 | see notes | Used to control firmware upgrades. | 10 | Bootloader
1034 | 0x40A | Reserved Flags | uint16 | - | Reserved for future use. Do not modify. | 10 | Bootloader
1036 | 0x40C | Bootloader Flash Address | uint32 | >= 0x20000 | Flash location to store firmware upgrades. Do not modify. Locations < FLASH_SIZE write to internal flash, locations > FLASH_SIZE write to external flash (if available). FLASH_SIZE depends on MCU_ID (35 --> 0x40000). | 10 | Bootloader
1034 | 0x40A
1040 | 0x410 | RESERVED | uint16_t | - | Reserved for future use. | 10 | Bootloader
2000 | 0x7D0 | Low Battery Threshold | float | see notes | Low Battery Threshold | 10
2002 | 0x7D2
2004 | 0x7D4 | Dead Battery Threshold | float | see notes | Dead Battery Threshold | 10
2006 | 0x7D6

## Base Station Config
LORD Sensing Base Station EEPROM Map
EEPROM Address |  | Option | Type | Valid Values | Description | Min FW Version | Category
decimal | hex
20 | 0x14 | Serial ID | uint32 |  | The serial id for the device. Combine with the model number and model option for the full Serial Number. | 1 | -
22 | 0x16
40 | 0x28 | Beacon Configuration | uint16 | [0-3] | Controls beacon echos over the port (0 = off, 1 = echo rx, 2 = echo tx, 3 = echo both) | 3.34 | beacon
46 | 0x2E | Model Number | uint16 |  | The model number of the device. | 3 | -
48 | 0x30 | Model Option | uint16 |  | The model option of the device. | 3 | -
50 | 0x32 | Base Station Address | uint16 | 4660 | The address of the device. This is the same for all Base Stations. Do Not Change. | 1 | -
90 | 0x5A | Radio Frequency | uint16 | see notes | The radio frequency of the device. Nodes on this same frequency can be communicated with. | 1 | -
94 | 0x5E | Transmit Power Level | uint16 | see notes | The transmit output power of the radio. | 1 | -
96 | 0x60 | Beacon Source | uint16 | see notes | The source of the beacon clock. | 3 | beacon
100 | 0x64 | PPS Edge Detection | uint16 | see notes | Tells the microcontroller which edge to interrupt off of when using a PPS for timing. | 1 | beacon
108 | 0x6C | Firmware Version (1) | uint16 | see notes | The firmware version of the device (Part 1). | 4 | -
110 | 0x6E | Firmware Version (2) | uint16 | see notes | The firmware version of the device (Part 2). | 4 | -
112 | 0x70 | (LEGACY) Model Number | uint16 |  | (LEGACY, Replaced by EEPROMs 46 + 48) The model number of the device. | 1 | -
114 | 0x72 | (LEGACY) Serial ID | uint16 |  | (LEGACY, Replaced by EEPROM 20) The serial id for the device. | 1 | -
118 | 0x76 | Radio ID | uint16 |  | The id of the radio on the device. | 1 | -
120 | 0x78 | Microcontroller ID | uint16 | see notes | The id of the microcontroller on the device. | 1 | -
122 | 0x7A | Firmware Arch Version | uint16 |  | The architecture version of the device used to determine which version of firmware can be applied to it. | 3 | -
124 | 0x7C | ASPP Version LXRS | uint16 |  | The ASPP version (msb.lsb) the device supports when in LXRS mode. | 4.0 | -
126 | 0x7E | ASPP Version LXRS+ | uint16 |  | The ASPP version (msb.lsb) the device supports when in LXRS+ mode. | 5.0 | -
128 | 0x80 | Analog Port 1 - Node Address | uint16 |  | The node address to associate with analog out, port 1. | 3 | analog
130 | 0x82 | Analog Port 1 - Channel | uint16 |  | The node channel number to associate with analog out, port 1. | 3 | analog
132 | 0x84 | Analog Port 1 - Max Float | float |  | The maximum float value if converting from float data with analog out, port 1. | 3 | analog
134 | 0x86
136 | 0x88 | Analog Port 1 - Min Float | float |  | The minimum float value if converting from float data with analog out, port 1. | 3 | analog
138 | 0x8A
140 | 0x8C | Analog Port 2 - Node Address | uint16 |  | The node address to associate with analog out, port 2. | 3 | analog
142 | 0x8E | Analog Port 2 - Channel | uint16 |  | The node channel number to associate with analog out, port 2. | 3 | analog
144 | 0x90 | Analog Port 2 - Max Float | float |  | The maximum float value if converting from float data with analog out, port 2. | 3 | analog
146 | 0x92
148 | 0x94 | Analog Port 2 - Min Float | float |  | The minimum float value if converting from float data with analog out, port 2. | 3 | analog
150 | 0x96
152 | 0x98 | Analog Port 3 - Node Address | uint16 |  | The node address to associate with analog out, port 3. | 3 | analog
154 | 0x9A | Analog Port 3 - Channel | uint16 |  | The node channel number to associate with analog out, port 3. | 3 | analog
156 | 0x9C | Analog Port 3 - Max Float | float |  | The maximum float value if converting from float data with analog out, port 3. | 3 | analog
158 | 0x9E
160 | 0xA0 | Analog Port 3 - Min Float | float |  | The minimum float value if converting from float data with analog out, port 3. | 3 | analog
162 | 0xA2
164 | 0xA4 | Analog Port 4 - Node Address | uint16 |  | The node address to associate with analog out, port 4. | 3 | analog
166 | 0xA6 | Analog Port 4 - Channel | uint16 |  | The node channel number to associate with analog out, port 4. | 3 | analog
168 | 0xA8 | Analog Port 4 - Max Float | float |  | The maximum float value if converting from float data with analog out, port 4. | 3 | analog
170 | 0xAA
172 | 0xAC | Analog Port 4 - Min Float | float |  | The minimum float value if converting from float data with analog out, port 4. | 3 | analog
174 | 0xAE
176 | 0xB0 | Analog Port 5 - Node Address | uint16 |  | The node address to associate with analog out, port 5. | 3 | analog
178 | 0xB2 | Analog Port 5 - Channel | uint16 |  | The node channel number to associate with analog out, port 5. | 3 | analog
180 | 0xB4 | Analog Port 5 - Max Float | float |  | The maximum float value if converting from float data with analog out, port 5. | 3 | analog
182 | 0xB6
184 | 0xB8 | Analog Port 5 - Min Float | float |  | The minimum float value if converting from float data with analog out, port 5. | 3 | analog
186 | 0xBA
188 | 0xBC | Analog Port 6 - Node Address | uint16 |  | The node address to associate with analog out, port 6. | 3 | analog
190 | 0xBE | Analog Port 6 - Channel | uint16 |  | The node channel number to associate with analog out, port 6. | 3 | analog
192 | 0xC0 | Analog Port 6 - Max Float | float |  | The maximum float value if converting from float data with analog out, port 6. | 3 | analog
194 | 0xC2
196 | 0xC4 | Analog Port 6 - Min Float | float |  | The minimum float value if converting from float data with analog out, port 6. | 3 | analog
198 | 0xC6
200 | 0xC8 | Analog Port 7 - Node Address | uint16 |  | The node address to associate with analog out, port 7. | 3 | analog
202 | 0xCA | Analog Port 7 - Channel | uint16 |  | The node channel number to associate with analog out, port 7. | 3 | analog
204 | 0xCC | Analog Port 7 - Max Float | float |  | The maximum float value if converting from float data with analog out, port 7. | 3 | analog
206 | 0xCE
208 | 0xD0 | Analog Port 7 - Min Float | float |  | The minimum float value if converting from float data with analog out, port 7. | 3 | analog
210 | 0xD2
212 | 0xD4 | Analog Port 8 - Node Address | uint16 |  | The node address to associate with analog out, port 8. | 3 | analog
214 | 0xD6 | Analog Port 8 - Channel | uint16 |  | The node channel number to associate with analog out, port 8. | 3 | analog
216 | 0xD8 | Analog Port 8 - Max Float | float |  | The maximum float value if converting from float data with analog out, port 8. | 3 | analog
218 | 0xDA
220 | 0xDC | Analog Port 8 - Min Float | float |  | The minimum float value if converting from float data with analog out, port 8. | 3 | analog
222 | 0xDE
224 | 0xE0 | Analog Out Enable | uint16 | [0, 1] | Whether analog pairing is enabled (1) or disabled (0). | 3 | analog
226 | 0xE2 | Analog Out Timeout - Time | uint16 |  | The number of seconds before an analog out port times out (0 disables timeout). | 3 | analog
228 | 0xE4 | Analog Out Timeout - Voltage | float | [0.0-3.0] | The voltage that the port will go to once it times out. | 3 | analog
230 | 0xE6
232 | 0xE8 | Button 1 - Long Press - Function | uint16 | see notes | The function to use when button 1 is long-pressed. | 3 | button
234 | 0xEA | Button 1 - Long Press - Node | uint16 |  | The node address associated with button 1, long press. | 3 | button
236 | 0xEC | Button 1 - Short Press - Function | uint16 | see notes | The function to use when button 1 is short-pressed. | 3 | button
238 | 0xEE | LED Action | uint16 |  | Controls the LED action of the device. | 2.27 | -
242 | 0xF2 | Reset Counter | uint16 |  | Records the number of times the device has been reset (both soft reset and power cycling). | 1 | -
250 | 0xFA | Cycle Power | uint16 | [1, 2] | Used to cycle the power on the device. (1-power, 2-radio) | 1 | -
256 | 0x100 | Button 1 - Short Press - Node | uint16 |  | The node address associated with button 1, short press. | 3 | button
258 | 0x102 | Button 2 - Long Press - Function | uint16 | see notes | The function to use when button 2 is long-pressed. | 3 | button
260 | 0x104 | Button 2 - Long Press - Node | uint16 |  | The node address associated with button 2, long press. | 3 | button
262 | 0x106 | Button 2 - Short Press - Function | uint16 | see notes | The function to use when button 2 is short-pressed. | 3 | button
264 | 0x108 | Button 2 - Short Press - Node | uint16 |  | The node address associated with button 2, short press. | 3 | button
280 | 0x118 | Region Code | uint16 | see notes | The region the device will be operated in. | 3 | -
348 | 0x15C | Analog Exceedance - Max | float |  | The maximum value for the analog exceedance setting. | 3 | analog
350 | 0x15E
352 | 0x160 | Analog Exceedance - Min | float |  | The minimum value for the analog exceedance setting. | 3 | analog
354 | 0x162
356 | 0x164 | Analog Exceedance Enable | uint16 | [0, 1] | Whether analog exceedance is enabled (1) or disabled (0). | 3 | analog
480 | 0x1E0 | Min software version required (Major) | uint16 |  | The minimum SensorConnect version required to support the BaseStation. (major) | - | -
482 | 0x1E2 | Min software version required (Minor) | uint16 |  | The minimum SensorConnect version required to support the BaseStation. (minor) | - | -
1024 | 0x400 | Communication Protocol | uint16 | [0, 1] | The communication protocol the device is currently set for. (0 = LXRS, 1 = LXRS+) | 5.0 | -
1030 | 0x406 | Bootloader version | uint16 | [0-65535] | Bootloader version # (continuous). If changed, the bootloader will rewrite it next time the node boots. | - | Bootloader
1032 | 0x408 | Firmware upgrade flag(s) | uint16 | see notes | Used to control firmware upgrades. | - | Bootloader
1034 | 0x40A | Reserved Flags | uint16 | - | Reserved for future use. Do not modify. | - | Bootloader
1036 | 0x40C | Bootloader Flash Address | uint32 | >= 0x20000 | Flash location to store firmware upgrades. Do not modify. Locations < FLASH_SIZE write to internal flash, locations > FLASH_SIZE write to external flash (if available). FLASH_SIZE depends on MCU_ID (35 --> 0x40000). | - | Bootloader
1034 | 0x40A
1040 | 0x410 | RESERVED | uint16_t | - | Reserved for future use. | - | Bootloader

## Config Values
Microcontroller ID
The ID of the Microcontroller on the device.
Value | ID
31 | 18F452, 20MHz
32 | 18F4620, 20MHz
33 | 18F46K20, 40MHz
34 | 18F67K90, 40MHz
35 | EFM32WG990F256, 48MHz
Data Collection Method
The data collection method when performing Sync or Async Sampling.
Value | ID | Details
1 | Log Only | Data is logged to Node's internal memory to be downloaded later.
2 | Transmit Only | Data is transmitted wirelessly over the air.
3 | Log and Transmit | Data is both logged to the Node and transmitted over the air.
Packet Data Type
Value transmitted in data packets to determine how the data should be parsed.
Value | ID | Details
1 | uint16 (bit-shifted) | Value needs to be divided by 2 after receiving
2 | float
3 | uint16 (12-bit resolution)
4 | uint32
5 | uint16 (odd average) | mean average where the # of samples used is an odd power of 2
6 | uint16 (even average) | mean average where the # of samples used is an even power of 2
7 | uint16 (full 16-bit resolution)
8 | float (no cals applied) | 4-byte float, with no calibration applied
9 | uint24
10 | uint16 (18-bit) | uint16 value from an 18-bit device, truncated, (val << 2) to get correct value
11 | int24 (20-bit) | 24-bit signed value from a 20-bit device
12 | int16 (20-bit) | 16-bit signed value from a 20-bit device, (val << 4) to get correct value
Data Format
The type of data formats that can be applied to a Wireless Node.
Value | ID | Details
1 | raw uint16 | no calibration will be applied
2 | calibrated float | cal coefficients will be applied
3 | raw uint24 (18-bit) | no calibration will be applied, (uint18 range)
4 | raw int24 (20-bit) | no calibration will be applied, (int20 range)
5 | raw int16 | no calibration will be applied
Sampling Mode
The sampling mode that the device's configuration is applied for.
Value | ID | Details
1 | Sync | Synchronized Sampling mode.
2 | Sync Burst | Synchronized Sampling, Burst mode.
3 | Async | Asynchronous Sampling mode (LDC).
4 | Armed Datalogging | Armed Datalogging Sampling mode.
Sync Sampling Configuration
Configures the Sync Sampling mode of operation.
Value | ID | Details
29696 | Continuous | Data will be transmitted in standard, non-burst synchronized sampling mode.
62976 | Burst | Data will be transmitted in a burst synchronized sampling mode.
Default Mode
The mode a Wireless Node starts in, and goes into after a user inactivity timeout.
Value | ID | Details
0 | Idle Mode | Idle state in which the Node can be communicated with normally. (default)
1 | Async Sampling | Starts Async Sampling (LDC)
4 | Armed Datalogging | Starts datalogging (but with no timestamp).
5 | Sleep | Low-Power Sleep Mode (must be set to idle to communicate).
6 | Sync Sampling | Starts Sync Sampling (must hear a beacon to start sampling).
Radio Frequency
The Wireless Frequency of a device.
Value | ID | Details
11 | Frequency 11 | 2.405 Ghz
12 | Frequency 12 | 2.410 Ghz
13 | Frequency 13 | 2.415 Ghz
14 | Frequency 14 | 2.420 Ghz
15 | Frequency 15 | 2.425 Ghz
16 | Frequency 16 | 2.430 Ghz
17 | Frequency 17 | 2.435 Ghz
18 | Frequency 18 | 2.440 Ghz
19 | Frequency 19 | 2.445 Ghz
20 | Frequency 20 | 2.450 Ghz
21 | Frequency 21 | 2.455 Ghz
22 | Frequency 22 | 2.460 Ghz
23 | Frequency 23 | 2.465 Ghz
24 | Frequency 24 | 2.470 Ghz
25 | Frequency 25 | 2.475 Ghz
26 | Frequency 26 | 2.480 Ghz
Transmit Power
The transmit power level of a device.
For Nodes with FW >= 10.0 and BaseStations with FW >= 4.0:
Value | ID | Details
20 | 20 dBm | 100 mw
16 | 16 dBm | 39 mw
10 | 10 dBm | 10 mw
5 | 5 dBm | 3 mw
0 | 0 dBm | 1 mw
For Nodes with FW < 10.0 and BaseStations with FW < 4.0:
Value | ID | Details
25619 | 16 dBm | 39 mw
25615 | 10 dBm | 10 mw
25611 | 5 dBm | 3 mw
25607 | 0 dBm | 1 mw
Retransmission
Whether retransmission (used for lossless Sync Sampling) is on or off.
Value | ID | Details
0 | Off | Retransmission is turned off.
1 | On | Retransmission is turned on
2 | Disabled | Software only lock. Software should not attempt to enable Retransmission.
Trigger Type
The type of a trigger that caused a datalogging session.
Value | ID | Details
0 | User Init | The user started this datalogging session manually by sending a start sampling command.
1 | Ceiling | The trigger was caused by a ceiling event.
2 | Floor | The trigger was caused by a floor event.
3 | Ramp Up | The trigger was caused by a ramp-up event.
4 | Ramp Down | The trigger was caused by a ramp-down event.
Settling Time
The settling time used for thermocouple and voltage inputs.
Value | ID | Details
1 | 4 ms | fastest settling
2 | 8 ms
3 | 16 ms
4 | 32 ms
5 | 40 ms
6 | 48 ms
7 | 60 ms
8 | 101 ms | 90db [60Hz Rejection]
9 | 120 ms | 80db [50Hz Rejection]
10 | 120 millisecond settling time | 65db [50+60Hz Rejection]
11 | 160 millisecond settling time | 69db [50+60Hz Rejection]
12 | 200 millisecond settling time | highest resolution
Thermocouple Type
The Thermocouple Type of a Wireless thermocouple channel.
Value | ID | Details
0 | Uncompensated | no thermocouple type
1 | K Type
2 | J Type
3 | R Type
4 | S Type
5 | T Type
6 | E Type
7 | B Type
8 | N Type
9 | Custom Polynomial
Armed Datalogging Sample Rate
The Sample Rate for Armed Datalogging sampling mode.
Value | ID
71 | 10 kHz
70 | 9 kHz
69 | 8 kHz
68 | 7 kHz
67 | 6 kHz
66 | 5 kHz
65 | 4 kHz
64 | 3 kHz
63 | 2 kHz
62 | 1 kHz
0 | 4096 Hz
1 | 2048 Hz
2 | 1024 Hz
3 | 512 Hz
4 | 256 Hz
5 | 128 Hz
6 | 64 Hz
7 | 32 Hz
Sample Rate
The Sample Rate (device and sampling mode dependent).
Value | ID
60 | 104170 Hz
58 | 78125 Hz
57 | 62500 Hz
56 | 25000 Hz
55 | 12500 Hz
49 | 3200 Hz
48 | 1600 Hz
47 | 800 Hz
46 | 300 Hz
62 | 1 kHz
63 | 2 kHz
64 | 3 kHz
65 | 4 kHz
66 | 5 kHz
67 | 6 kHz
68 | 7 kHz
69 | 8 kHz
70 | 9 kHz
71 | 10 kHz
72 | 20 kHz
73 | 30 kHz
74 | 40 kHz
75 | 50 kHz
76 | 60 kHz
77 | 70 kHz
78 | 80 kHz
79 | 90 kHz
80 | 100 kHz
98 | 910 Hz
100 | 8192 Hz
101 | 4096 Hz
102 | 2048 Hz
103 | 1024 Hz
104 | 512 Hz
105 | 256 Hz
106 | 128 Hz
107 | 64 Hz
108 | 32 Hz
109 | 16 Hz
110 | 8 Hz
111 | 4 Hz
112 | 2 Hz
113 | 1 Hz
114 | 1 sample every 2 seconds
115 | 1 sample every 5 seconds
116 | 1 sample every 10 seconds
117 | 1 sample every 30 seconds
118 | 1 sample every 1 minute
119 | 1 sample every 2 minutes
120 | 1 sample every 5 minutes
121 | 1 sample every 10 minutes
122 | 1 sample every 30 minutes
123 | 1 sample every 60 minutes
127 | 1 sample every 24 hours
Region Code
The region that the device will be operated in.
Value | ID
0 | USA
1 | Europe
2 | Japan
3 | Other
4 | Brazil
Beacon Source
The source of a Base Station's beacon.
Value | ID
0 | None - No Beacon
1 | Internal Timer
2 | Internal PPS
3 | External PPS
PPS Edge Detection
Which edge to interrupt off of when using a PPS for timing.
For BaseStations with FW >= 4.0:
Value | ID
0 | Off
1 | Rising Edge
2 | Falling Edge
3 | Rising and Falling Edge
For BaseStations with FW < 4.0:
Value | ID
0 | Falling Edge
1 | Rising Edge
Button Actions
The actions that can be performed by a Base Station's physical buttons.
Value | ID | Details
0 | Node Sleep | Puts a Node into sleep mode.
1 | Node Set to Idle | Puts a Node into idle mode.
2 | Enable Beacon | Enable the Base Station's beacon.
3 | Disable Beacon | Disables the Base Station's beacon.
4 | Start Node Non-Sync Sampling | Starts a Node sampling in Non-Sync mode.
5 | Start Node Sync Sampling | Starts a Node sampling in Sync mode.
6 | Start Node Armed Datalogging | Starts a Node sampling in Armed Datalogging mode.
7 | Cycle Power | Cycles the Base Station's power.
65535 | Disabled | Disables the button functionality.
Firmware Version
The firmware version of the device.
For Nodes with FW >= 10.0 and BaseStations with FW >= 4.0:
The firmware version of Base Stations and Wireless Nodes are repesented in the following format:
Major.Revision
where Major is the first byte (MSB) of EEPROM 108, and Revision is a 3 byte value created by combining the last byte (LSB) of EEPROM 108 and both bytes of EEPROM 110.
For example:
uint16_t fwValue1 = readEeprom(108);
uint16_t fwValue2 = readEeprom(110);
uint8_t msb1 = fwValue1 >> 8;
uint8_t lsb1 = fwValue1 & 0x00FF;
uint8_t msb2 = fwValue2 >> 8;
uint8_t lsb2 = fwValue2 & 0x00FF;
uint16_t hiWord = (0 << 8) ¦ lsb1;
uint16_t loWord = (msb2 << 8) ¦ lsb2;
uint8_t major = msb1;
uint32_t revision = static_cast<uint32_t>(hiWord) << 16 ¦ static_cast<uint32_t>(loWord)
For Nodes with FW < 10.0 BaseStations with FW < 4.0:
The firmware version of Base Stations and Wireless Nodes are repesented in the following format:
Major.Minor
where Major is the first byte (MSB) of EEPROM 108, and Minor is the last byte (LSB) of EEPROM 108.
For example:
uint16_t fwValue = readEeprom(108);
uint8_t major = fwValue >> 8;
uint8_t minor = fwValue & 0x00FF;
Channel Mask
The Wireless Nodes use a channel mask to determine which channels are active and inactive on the device.
Channels that are active will be sampled during any type of sampling mode. The mask is a 16-bit value. Each of the 16 bits of the mask correspone to one of the Node's channels (if that Node supports that channel). Bit 1 correpsonds to channel 1, bit 2 to channel 2, bit 8 to channel 8, etc. When the bit is set to 1, the channel is active. When the bit is set to 0, the channel is inactive.
For example:
uint16_t ch1_ch6 = 33; //0000 0000 0010 0001 (Channel 1 and Channel 6 enabled)
uint16_t ch3thru8 = 252; //0000 0000 1111 1100 (Channel 3 through 8 enabled)
Sampling Delay
The sampling delay sets the duration (in milliseconds) of pulsed power sensor excitation i.e. the amount of time between sensor excitation power up and A/D sampling. Low-impedance sensors with slow rise times can exhibit longer settling time, and thus benefit from a larger delay before sampling. Smaller delay values reduce power consumption during LDC. Note that this value only applies to sample rates less than 32 Hz, since sensor excitation power is applied full time for sample rates of 32 Hz and above. A value of 10,000 may be entered into this location to provide full time excitation power for sample rates below 32 Hz.
Set to Idle Interval
How often the Node wakes up and listens for a 'set to idle' command.
The following equation can be used to determine the specific value needed for a wake interval defined in seconds:
Value = (7680 / Desired_Interval_in_seconds)
For example, if the user wants the node to check for a wake command every 10 seconds, then this location would be programmed with a value of
(7680/10) = 768
The minimum wake interval is internally limited to 1 second (a value of 7680), and the maximum wake interval is limited to approximately 15 seconds (a value of 512). Any values outside of this range will be automatically truncated.
NOTE: After modifying this value, the node must be reset in order for the changes to take effect.
Calibration Coefficients
Calibration Coefficients are linear scaling filters applied to convert a channel's raw bit value to physical or engineering units. Please see the individual Node's hardware manual for more information on general calibration procedures.
Note: The calibration coefficients slope and offset is different from a Node's hardware gain and offset settings.
Note: When a Node is transmitting 2-byte integer, uncalibrated values, Calibration Coefficients represent a post-processing step. The conversion from bits to physical units needs to take place after the data is received, and not on the Node. However, when the Node is transmitting 4-byte float, calibrated value, the conversion will occur on the Node and the received data values will already have been converted.
Each Channel requires 10 bytes to hold the equation, unit, slope, and offset values:
Description | Type
Equation ID | uint8
Unit ID | uint8
Slope | float
Offset | float
Cal Coefficients - Equation Type
The equation type for cal coefficients.
Value | ID | Details
0 | y = x | value = bits
4 | y = mx + b | value = (slope * bits) + offset
Cal Coefficients - Unit
The unit type for cal coefficients.
Value | ID | Details
0 | unknown | no unit or unknown unit
1 | bits | raw bits
2 | Strain
3 | microStrain
4 | acceleration due to gravity | G's
5 | meters per second squared
6 | volts
7 | milliVolts
8 | microVolts
9 | degrees Celsius
10 | Kelvin
11 | degrees Fahrenheit
12 | meters
13 | millimeters
14 | micrometers
15 | pound force
16 | Newtons
17 | kiloNewtons
18 | kilograms
19 | bar
20 | pounds per square inch
21 | atmospheric pressure
22 | millimeters of mercury
23 | Pascal
24 | megaPascal
25 | kiloPascal
26 | degrees
27 | degrees per second
28 | radians per second
29 | percent
30 | revolutions per minute
31 | hertz
32 | percent relative humidity
33 | milliVolt/Volt
34 | milli-G's
35 | feet per second squared
36 | percent life
37 | count
38 | feet
39 | inches
40 | yards
41 | miles
42 | nautical miles
43 | thousandths of an inch
44 | hundredths of an inch
45 | kilometers
46 | centimeters
47 | watts per square meter
48 | microEinstein
49 | pound
50 | watt
51 | milliwatt
52 | horsepower
53 | Volt-Ampere Reactive
54 | watt-hour
55 | kilowatt-hour
56 | Volt-Ampere Reactive Hour
57 | Kilovolt-Ampere Reactive Hour
58 | Ampere
59 | Milliampere
60 | Microampere
61 | MilliBar
62 | Inches of Mercury
63 | dBm
64 | Kilohertz
65 | Radians
66 | Meters per Second
67 | Kilometers per Second
68 | Kilometers per Hour
69 | Miles per Hour
70 | Knots
71 | Cubic Meter
72 | Cubic Feet
73 | Liters
74 | Gallon
75 | Cubic Meters Per Second
76 | Cubic Feet Per Second
77 | Newton Meter
78 | Foot Pounds
79 | Inch Pounds
80 | Seconds
81 | Nanoseconds
82 | Microseconds
83 | Milliseconds
84 | Minutes
85 | Hours
86 | Days
87 | Weeks
88 | value
89 | Gauss
90 | G-Seconds
91 | Seconds per Second | (clock drift)
92 | dBHz
93 | kg Per Meter cubed
94 | unitless
Upgrade Flags
Controls firmware upgrade behavior - Bootloader version 4
In general, these flags are handled by the node firmware. Except to set the override or open comm flags, this setting should be left alone.
Flag | Description
0x01 - Upgrade | Causes bootloader to perform an upgrade. Will only succeed if a valid & compatible image exists in flash. See bootload address for the flash location. Compatible images must have the same model number, firmware architecture, and MCU ID.
0x02 - Override | Skips checking the new firmware for compatibility.
0x04 - Open Comm | Opens a communications channel. Generally this is only supported on base stations.
0x80 - Testing | Set by the bootloader after an upgrade is performed. If the upgrade is succesful, the flag is immediately cleared by the new firmware. The flag exists to increase reliability and to prevent excessive wear on the flash memory. DO NOT SET THIS FLAG.
Bootloader Flash Address
Determines were firmware upgrades are stored in flash memory.
This is handled internally by the node firmware. User modification is not supported at this time.
Values less than FLASH_SIZE refer to internal flash.
Values greater than FLASH_SIZE refer to external flash.
FLASH_SIZE depends on MCU_ID. For an MCU_ID of 35, FLASH_SIZE is 0x40000.
Values less than BOOTLOADER_SIZE (0x6000 for bootloader version 4) are invalid.
Flash ID
Identifier for the Flash Chip
Value | Model | Total Size (MiB) | Block Size (KiB) | Page Size (Bytes)
0 | SST25VF032B | 4 | 64 | 256
1 | IS25WP128 | 16 | 64 | 256
Battery Thresholds
V-Link-200: Values are user adjustable: Low Battery Threshold: EEPROM Addr = 2000; Default Value = 10.5 Dead Battery Threshold: EEPROM Addr = 2004; Default Value = 6.3V G-Link-200: Values are not user adjustable. Also the G-Link-200 uses a boost converter which makes the low battery conditions harder to pick-up consistently. Low Battery Threshold: Default Value = 2.8V Dead Battery Threshold: Default Value = 2.0V